Standard Operating Procedure: Equipment Calibration

1.	Purpose

This Standard Operating Procedure establishes the process for calibrating and maintaining measurement and test equipment to ensure accuracy and compliance with NYS Office of Cannabis Management (OCM) GMP requirements.

2.	Scope

This SOP applies to all measurement and test equipment used in production, quality control, and environmental monitoring at licensed cannabis facilities in New York State.

3.	Prerequisites

• Training on calibration procedures
• Understanding of measurement principles
• Knowledge of equipment specifications
• Access to calibration standards
• Familiarity with documentation requirements

4.	Responsibilities

Quality Assurance Manager:
• Oversee calibration program
• Approve calibration procedures
• Review out-of-tolerance events
• Ensure compliance

Calibration Technician:
• Perform calibrations
• Maintain calibration standards
• Document calibration activities
• Report out-of-tolerance conditions

Equipment Users:
• Check calibration status before use
• Report equipment issues
• Do not use out-of-calibration equipment
• Handle equipment properly

Facilities/Maintenance:
• Support calibration activities
• Maintain equipment in good condition
• Coordinate external calibrations

5.	Materials

• Calibrated reference standards
• Calibration tools and accessories
• Calibration certificates
• Calibration labels/stickers
• Calibration log/database
• Equipment manuals
• Out-of-tolerance investigation forms
• Environmental monitoring equipment

6.	Procedure

6.1 Calibration Program Scope
• Equipment requiring calibration includes:
  - Scales and balances
  - Thermometers and temperature probes
  - Humidity sensors
  - pH meters
  - Conductivity/TDS meters
  - Pressure gauges
  - Flow meters
  - Timers
  - Environmental monitors
  - Laboratory instruments

6.2 Calibration Schedule
• Establish calibration frequency based on:
  - Manufacturer recommendations
  - Equipment criticality
  - Historical stability
  - Regulatory requirements
• Typical frequencies:
  - Scales/balances: Annually + daily verification
  - Temperature devices: Annually
  - pH meters: Before each use + annually
  - Environmental monitors: Quarterly
• Maintain calibration schedule/calendar
• Track upcoming calibrations monthly

6.3 Calibration Standards
• Use traceable reference standards
• Standards must be more accurate than equipment (4:1 ratio preferred)
• Maintain calibration certificates for standards
• Store standards per manufacturer requirements
• Calibrate standards at appropriate intervals
• Replace standards before expiration

6.4 Calibration Procedure
• Allow equipment to stabilize at ambient conditions
• Verify calibration standards are in tolerance
• Clean equipment before calibration
• Perform calibration per equipment SOP:
  - Multiple points across operating range
  - Ascending and descending readings
  - Record all measurements
• Compare readings to acceptance criteria
• Adjust if necessary and re-verify
• Document all results

6.5 Calibration Acceptance Criteria
• Define tolerance for each equipment type:
  - Scales: ± 0.1% of reading or as specified
  - Thermometers: ± 0.5°C or as specified
  - pH meters: ± 0.1 pH units
  - Humidity sensors: ± 3% RH
• Document criteria in equipment-specific SOPs
• Adjust tolerances based on use requirements

6.6 Calibration Documentation
• Record for each calibration:
  - Equipment identification
  - Calibration date
  - Technician name
  - Standards used (with traceability)
  - As-found readings
  - Adjustments made
  - As-left readings
  - Pass/fail determination
  - Next calibration due date
• Maintain calibration certificates
• Retain records per document control SOP

6.7 Calibration Labels
• Apply calibration label to equipment showing:
  - Calibration date
  - Next due date
  - Technician initials
  - Status (passed/limited use)
• Remove labels from failed equipment
• Use color-coded labels by year if desired

6.8 Out-of-Tolerance Handling
• When equipment found out of tolerance:
  - Remove from service immediately
  - Label as "Out of Calibration"
  - Investigate impact on products
  - Review data since last calibration
  - Determine if products affected
  - Document investigation
  - Implement corrective actions
• Assess whether equipment can be adjusted or must be replaced

6.9 Verification vs. Calibration
• Daily verification: Quick check against known standard
  - Confirms continued accuracy
  - Does not replace calibration
  - Document verification results
• Calibration: Full procedure against traceable standards
  - Adjusts equipment if needed
  - Generates calibration certificate

6.10 External Calibration Services
• Use accredited calibration laboratories
• Verify laboratory qualifications
• Review calibration certificates for:
  - Traceability statements
  - Measurement uncertainty
  - As-found/as-left data
  - Standards used
• Maintain copies of all certificates

7.	References

• NYS OCM GMP Guidance
• ISO 17025 Laboratory Accreditation
• NIST Traceability Requirements
• Source: https://cannabis.ny.gov/

8.	Definitions

Calibration: Comparison of a measurement device against a reference standard and adjustment if necessary to meet specifications.

Traceability: An unbroken chain of comparisons relating a measurement result to a national or international standard.

Tolerance: The acceptable range of variation from a specified value.

Reference Standard: A measurement device with known accuracy used to calibrate other equipment.

As-Found/As-Left: The condition of equipment before calibration (as-found) and after any adjustments (as-left).

Related SOPs and Cross-References

Category: MAINTENANCE - NYS OCM

The following SOPs are related to this procedure:

GEN-012: Preventive Maintenance - Equipment care

PROC-029: Quality Control Testing - Test equipment

GEN-001: Document Control - Calibration records

Note: This SOP is part of a comprehensive cannabis compliance system. SOPs marked "(OTHER STATES)" contain recommendations from other state programs and are suggestions, not NYS OCM requirements.